## Slide 1: Daily Monitoring Report Review
October 26, 2017

## Slide 2: PRAMP Daily Report
Internal Zero Span

## Slide 3: Something You Need To Know
Alberta Ambient Air Quality Objectives for SO2:
  1-hour average: 172ppb
  24-hour average: 48.0ppb.
THC background concentration:1.5 ppm.
THC ≈ CH4 + NMHC
NMHC 5 minute average ≥ 0.3ppm => Canister event
Wind Direction represents where the wind was blowing from.

## Slide 4: Something You Need To Know
Internal zero/span (IZS) test:
  is required to be run on a daily basis;
  shall not exceed 1 hour per calendar day;
  contains one zero point check and one span point check; both points check must achieve a minimum of 10 minutes stability; and the check result must meet calibration acceptance criteria.
 | zero drift | span drift | 
TRS | +/- 3 ppb Range: 0 - 100 ppb | +/- 10 % | Drift (%)
SO2 | +/- 15 ppb Range: 0- 500 ppb | +/- 10 % | 
THC/CH4/NMHC | +/- 1.2ppm (THC) +/- 0.6ppm (CH4/NMHC) Range: THC/CH4/NMHC: 40/20/20 ppm | +/- 10 % | 

## Slide 5
 | Reviews
TRS | Review Maximum, Minimum, and 24-hour average values Review data trend Hourly data recorded outside +/-3 ppb range (IZS zero requirement) during normal monitoring condition => analyzer failure
SO2 | Review Maximum, Minimum, and 24-hour average values Review data trend Check if there was any contravention (1-hr: 172ppb, 24-hr:48ppb)
THC/ CH4/ NMHC | Review Maximum, Minimum, and 24-hour average values Check if any THC < 1.5ppm was recorded THC ≈ CH4 + NMHC Check if NMHC > 0.1ppm or NMHCmax > 0.30ppm was recorded => canister event

## Slide 6: THC/CH4<1.5ppm TRS hourly data remains no change for hours

## Slide 7: NMHC >0.1ppm & NMHCmax >0.3ppm 5-min data (canister event)

## Slide 8
 | Reviews
Hourly max data | Check to see if there is abnormal A real event An abnormal spike
Chart: 
| 05:45 | 05:46 | 05:47 | 05:48 | 05:49 | 05:50 | 05:51 | 05:52 | 05:53 | 05:54 | 05:55 | 05:56 | 05:57 | 05:58 | 05:59 | 06:00 | 06:01 | 06:02 | 06:03 | 06:04 | 06:05 | 06:06 | 06:07 | 06:08 | 06:09 | 06:10 | 06:11 | 06:12 | 06:13 | 06:14 | 06:15 | 06:16 | 06:17 | 06:18 | 06:19 | 06:20 | 06:21 | 06:22 | 06:23 | 06:24 | 06:25 | 06:26 | 06:27 | 06:28 | 06:29 | 06:30 | 06:31 | 06:32 | 06:33 | 06:34 | 06:35 | 06:36 | 06:37 | 06:38 | 06:39 | 06:40 | 06:41 | 06:42 | 06:43 | 06:44 | 06:45 | 06:46 | 06:47 | 06:48 | 06:49 | 06:50 | 06:51 | 06:52 | 06:53 | 06:54 | 06:55 | 06:56 | 06:57 | 06:58 | 06:59 | 07:00 | 07:01 | 07:02 | 07:03 | 07:04 | 07:05 | 07:06 | 07:07 | 07:08 | 07:09 | 07:10 | 07:11 | 07:12 | 07:13 | 07:14 | 07:15 | 07:16 | 07:17 | 07:18 | 07:19 | 07:20 | 07:21 | 07:22 | 07:23 | 07:24 | 07:25 | 07:26 | 07:27 | 07:28 | 07:29 | 07:30 | 07:31 | 07:32 | 07:33 | 07:34 | 07:35 | 07:36 | 07:37 | 07:38 | 07:39 | 07:40 | 07:41 | 07:42 | 07:43 | 07:44 | 07:45 | 07:46 | 07:47 | 07:48 | 07:49 | 07:50 | 07:51 | 07:52 | 07:53 | 07:54 | 07:55 | 07:56 | 07:57 | 07:58 | 07:59 | 08:00 | 08:01 | 08:02 | 08:03 | 08:04 | 08:05 | 08:06 | 08:07 | 08:08 | 08:09 | 08:10 | 08:11 | 08:12 | 08:13 | 08:14 | 08:15 | 08:16 | 08:17 | 08:18 | 08:19 | 08:20 | 08:21 | 08:22 | 08:23 | 08:24 | 08:25 | 08:26 | 08:27 | 08:28 | 08:29 | 08:30 | 08:31 | 08:32 | 08:33 | 08:34 | 08:35 | 08:36 | 08:37 | 08:38 | 08:39 | 08:40 | 08:41 | 08:42 | 08:43 | 08:44 | 08:45 | 08:46 | 08:47 | 08:48 | 08:49 | 08:50 | 08:51 | 08:52 | 08:53 | 08:54 | 08:55 | 08:56 | 08:57 | 08:58 | 08:59 | 09:00 | 09:01 | 09:02 | 09:03 | 09:04 | 09:05 | 09:06 | 09:07 | 09:08 | 09:09 | 09:10 | 09:11 | 09:12 | 09:13 | 09:14 | 09:15 | 09:16 | 09:17 | 09:18 | 09:19 | 09:20 | 09:21 | 09:22 | 09:23 | 09:24 | 09:25 | 09:26 | 09:27 | 09:28 | 09:29 | 09:30 | 09:31 | 09:32 | 09:33 | 09:34 | 09:35 | 09:36 | 09:37 | 09:38 | 09:39 | 09:40 | 09:41 | 09:42 | 09:43 | 09:44 | 09:45 | 09:46 | 09:47 | 09:48 | 09:49 | 09:50 | 09:51 | 09:52 | 09:53 | 09:54 | 09:55 | 09:56 | 09:57 | 09:58 | 09:59 | 10:00 | 10:01 | 10:02 | 10:03 | 10:04 | 10:05 | 10:06 | 10:07 | 10:08 | 10:09 | 10:10 | 10:11 | 10:12 | 10:13 | 10:14 | 10:15 | 10:16 | 10:17 | 10:18 | 10:19 | 10:20 | 10:21 | 10:22 | 10:23 | 10:24 | 10:25 | 10:26 | 10:27 | 10:28 | 10:29 | 10:30 | 10:31 | 10:32 | 10:33 | 10:34 | 10:35 | 10:36 | 10:37 | 10:38 | 10:39 | 10:40 | 10:41 | 10:42 | 10:43 | 10:44 | 10:45 | 10:46 | 10:47 | 10:48 | 10:49 | 10:50 | 10:51 | 10:52 | 10:53 | 10:54 | 10:55 | 10:56 | 10:57 | 10:58 | 10:59 | 11:00 | 11:01 | 11:02 | 11:03 | 11:04 | 11:05 | 11:06 | 11:07 | 11:08 | 11:09 | 11:10 | 11:11 | 11:12 | 11:13 | 11:14 | 11:15 | 11:16 | 11:17 | 11:18 | 11:19 | 11:20 | 11:21 | 11:22 | 11:23 | 11:24 | 11:25 | 11:26 | 11:27 | 11:28 | 11:29 | 11:30 | 11:31 | 11:32 | 11:33 | 11:34 | 11:35 | 11:36 | 11:37 | 11:38 | 11:39 | 11:40 | 11:41 | 11:42 | 11:43 | 11:44 | 11:45 | 11:46 | 11:47 | 11:48 | 11:49 | 11:50 | 11:51 | 11:52 | 11:53 | 11:54 | 11:55 | 11:56 | 11:57 | 11:58 | 11:59 | 12:00 | 12:01 | 12:02 | 12:03 | 12:04 | 12:05 | 12:06 | 12:07 | 12:08 | 12:09 | 12:10 | 12:11 | 12:12 | 12:13 | 12:14 | 12:15 | 12:16 | 12:17 | 12:18 | 12:19 | 12:20 | 12:21 | 12:22 | 12:23 | 12:24 | 12:25 | 12:26 | 12:27 | 12:28 | 12:29 | 12:30 | 12:31 | 12:32 | 12:33 | 12:34 | 12:35 | 12:36 | 12:37 | 12:38 | 12:39 | 12:40 | 12:41 | 12:42 | 12:43 | 12:44 | 12:45 | 12:46 | 12:47 | 12:48 | 12:49 | 12:50 | 12:51 | 12:52 | 12:53 | 12:54 | 12:55 | 12:56 | 12:57 | 12:58 | 12:59 | 13:00 | 13:01 | 13:02 | 13:03 | 13:04 | 13:05 | 13:06 | 13:07 | 13:08 | 13:09 | 13:10 | 13:11 | 13:12 | 13:13 | 13:14 | 13:15 | 13:16 | 13:17 | 13:18 | 13:19 | 13:20 | 13:21 | 13:22 | 13:23 | 13:24 | 13:25 | 13:26 | 13:27 | 13:28 | 13:29 | 13:30 | 13:31 | 13:32 | 13:33 | 13:34 | 13:35 | 13:36 | 13:37 | 13:38 | 13:39 | 13:40 | 13:41 | 13:42 | 13:43 | 13:44 | 13:45 | 13:46 | 13:47 | 13:48 | 13:49 | 13:50 | 13:51 | 13:52 | 13:53 | 13:54 | 13:55 | 13:56 | 13:57 | 13:58 | 13:59 | 14:00 | 14:01 | 14:02 | 14:03 | 14:04 | 14:05 | 14:06 | 14:07 | 14:08 | 14:09 | 14:10 | 14:11 | 14:12 | 14:13 | 14:14 | 14:15 | 14:16 | 14:17 | 14:18 | 14:19 | 14:20 | 14:21 | 14:22 | 14:23 | 14:24 | 14:25 | 14:26 | 14:27 | 14:28 | 14:29 | 14:30 | 14:31 | 14:32 | 14:33 | 14:34 | 14:35 | 14:36 | 14:37 | 14:38 | 14:39 | 14:40 | 14:41 | 14:42 | 14:43 | 14:44 | 14:45 | 14:46 | 14:47 | 14:48 | 14:49 | 14:50 | 14:51 | 14:52 | 14:53 | 14:54 | 14:55 | 14:56 | 14:57 | 14:58 | 14:59 | 15:00 | 15:01 | 15:02 | 15:03 | 15:04 | 15:05 | 15:06 | 15:07 | 15:08 | 15:09 | 15:10 | 15:11 | 15:12 | 15:13 | 15:14 | 15:15 | 15:16 | 15:17 | 15:18 | 15:19 | 15:20 | 15:21 | 15:22 | 15:23 | 15:24 | 15:25 | 15:26 | 15:27 | 15:28 | 15:29 | 15:30 | 15:31 | 15:32 | 15:33 | 15:34 | 15:35 | 15:36 | 15:37 | 15:38 | 15:39 | 15:40 | 15:41 | 15:42 | 15:43 | 15:44 | 15:45 | 15:46 | 15:47 | 15:48 | 15:49 | 15:50 | 15:51 | 15:52 | 15:53 | 15:54 | 15:55 | 15:56 | 15:57 | 15:58 | 15:59 | 16:00 | 16:01 | 16:02 | 16:03 | 16:04 | 16:05 | 16:06 | 16:07 | 16:08 | 16:09 | 16:10 | 16:11 | 16:12 | 16:13 | 16:14 | 16:15 | 16:16 | 16:17 | 16:18 | 16:19 | 16:20 | 16:21 | 16:22 | 16:23 | 16:24 | 16:25 | 16:26 | 16:27 | 16:28 | 16:29 | 16:30 | 16:31 | 16:32 | 16:33 | 16:34 | 16:35 | 16:36 | 16:37 | 16:38 | 16:39 | 16:40 | 16:41 | 16:42 | 16:43 | 16:44 | 16:45 | 16:46 | 16:47 | 16:48 | 16:49 | 16:50 | 16:51 | 16:52 | 16:53 | 16:54 | 16:55 | 16:56 | 16:57 | 16:58 | 16:59 | 17:00 | 17:01 | 17:02 | 17:03 | 17:04 | 17:05 | 17:06 | 17:07 | 17:08 | 17:09 | 17:10 | 17:11 | 17:12 | 17:13 | 17:14 | 17:15 | 17:16 | 17:17 | 17:18 | 17:19 | 17:20 | 17:21 | 17:22 | 17:23 | 17:24 | 17:25 | 17:26 | 17:27 | 17:28 | 17:29 | 17:30 | 17:31 | 17:32 | 17:33 | 17:34 | 17:35 | 17:36 | 17:37 | 17:38 | 17:39 | 17:40 | 17:41 | 17:42 | 17:43 | 17:44 | 17:45 | 17:46 | 17:47 | 17:48 | 17:49 | 17:50 | 17:51 | 17:52 | 17:53 | 17:54 | 17:55 | 17:56 | 17:57 | 17:58 | 17:59 | 18:00 | 18:01 | 18:02 | 18:03 | 18:04 | 18:05 | 18:06 | 18:07 | 18:08 | 18:09 | 18:10 | 18:11 | 18:12 | 18:13 | 18:14 | 18:15 | 18:16 | 18:17 | 18:18 | 18:19 | 18:20 | 18:21 | 18:22 | 18:23 | 18:24 | 18:25 | 18:26 | 18:27 | 18:28 | 18:29 | 18:30 | 18:31 | 18:32 | 18:33 | 18:34 | 18:35 | 18:36 | 18:37 | 18:38 | 18:39 | 18:40 | 18:41 | 18:42 | 18:43 | 18:44 | 18:45 | 18:46 | 18:47 | 18:48 | 18:49 | 18:50 | 18:51 | 18:52 | 18:53 | 18:54 | 18:55 | 18:56 | 18:57 | 18:58 | 18:59 | 19:00 | 19:01 | 19:02 | 19:03 | 19:04 | 19:05 | 19:06 | 19:07 | 19:08 | 19:09 | 19:10 | 19:11 | 19:12 | 19:13 | 19:14 | 19:15 | 19:16 | 19:17 | 19:18 | 19:19 | 19:20 | 19:21 | 19:22 | 19:23 | 19:24 | 19:25 | 19:26 | 19:27 | 19:28 | 19:29 | 19:30 | 19:31 | 19:32 | 19:33 | 19:34 | 19:35 | 19:36 | 19:37 | 19:38 | 19:39 | 19:40 | 19:41 | 19:42 | 19:43 | 19:44 | 19:45 | 19:46 | 19:47 | 19:48 | 19:49 | 19:50 | 19:51 | 19:52 | 19:53 | 19:54 | 19:55 | 19:56 | 19:57 | 19:58 | 19:59 | 20:00 | 20:01 | 20:02 | 20:03 | 20:04 | 20:05 | 20:06 | 20:07 | 20:08 | 20:09 | 20:10 | 20:11 | 20:12 | 20:13 | 20:14 | 20:15 | 20:16 | 20:17 | 20:18 | 20:19 | 20:20 | 20:21 | 20:22 | 20:23 | 20:24 | 20:25 | 20:26 | 20:27 | 20:28 | 20:29 | 20:30 | 20:31 | 20:32 | 20:33 | 20:34 | 20:35 | 20:36 | 20:37 | 20:38 | 20:39 | 20:40 | 20:41 | 20:42 | 20:43 | 20:44 | 20:45 | 20:46 | 20:47 | 20:48 | 20:49 | 20:50 | 20:51 | 20:52 | 20:53 | 20:54 | 20:55 | 20:56 | 20:57 | 20:58 | 20:59 | 21:00 | 21:01 | 21:02 | 21:03 | 21:04 | 21:05 | 21:06 | 21:07 | 21:08 | 21:09 | 21:10 | 21:11 | 21:12 | 21:13 | 21:14 | 21:15 | 21:16 | 21:17 | 21:18 | 21:19 | 21:20 | 21:21 | 21:22 | 21:23 | 21:24 | 21:25 | 21:26 | 21:27 | 21:28 | 21:29 | 21:30 | 21:31 | 21:32 | 21:33 | 21:34 | 21:35 | 21:36 | 21:37 | 21:38 | 21:39 | 21:40 | 21:41 | 21:42 | 21:43 | 21:44 | 21:45 | 21:46 | 21:47 | 21:48 | 21:49 | 21:50 | 21:51 | 21:52 | 21:53 | 21:54 | 21:55 | 21:56 | 21:57 | 21:58 | 21:59 | 22:00 | 22:01 | 22:02 | 22:03 | 22:04 | 22:05 | 22:06 | 22:07 | 22:08 | 22:09 | 22:10 | 22:11 | 22:12 | 22:13 | 22:14 | 22:15 | 22:16 | 22:17 | 22:18 | 22:19 | 22:20 | 22:21 | 22:22 | 22:23 | 22:24 | 22:25 | 22:26 | 22:27 | 22:28 | 22:29 | 22:30 | 22:31 | 22:32 | 22:33 | 22:34 | 22:35 | 22:36 | 22:37 | 22:38 | 22:39 | 22:40 | 22:41 | 22:42 | 22:43 | 22:44 | 22:45 | 22:46 | 22:47 | 22:48 | 22:49 | 22:50 | 22:51 | 22:52 | 22:53 | 22:54 | 22:55 | 22:56 | 22:57 | 22:58 | 22:59 | 23:00 | 23:01 | 23:02 | 23:03 | 23:04 | 23:05 | 23:06 | 23:07 | 23:08 | 23:09 | 23:10 | 23:11 | 23:12 | 23:13 | 23:14 | 23:15 | 23:16 | 23:17 | 23:18 | 23:19 | 23:20 | 23:21 | 23:22 | 23:23 | 23:24 | 23:25 | 23:26 | 23:27 | 23:28 | 23:29 | 23:30 | 23:31 | 23:32 | 23:33 | 23:34 | 23:35 | 23:36 | 23:37 | 23:38 | 23:39 | 23:40 | 23:41 | 23:42 | 23:43 | 23:44 | 23:45 | 23:46 | 23:47 | 23:48 | 23:49 | 23:50 | 23:51 | 23:52 | 23:53 | 23:54 | 23:55 | 23:56 | 23:57 | 23:58 | 23:59
946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 946 947 947 947 947 947 947 947 946 947 946 946 946 947 946 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 947 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948 948
5.2 3.1 6 3.9 5.3 4.4 4.1 4.3 4.5 4.5 4.2 4.4 5 5.7 6.1 3.6 5.6 4.6 3.5 5.2 3.3 4 3.5 3.3 3.6 3.2 3.6 4.2 3.5 3.9 4.5 4.1 3.2 4 4.3 3.3 4.1 4.4 3.3 4.3 4.5 3.1 4.7 4.5 4.5 4.7 5.2 4.8 5 5.5 6.3 2.9 4.5 6.5 6 8.6 6.8 6.2 5.8 5.5 5.9 6.8 5 8 6 4.4 4.4 5.2 4.8 9.4 6.7 5.2 5.8 8.8 5.2 7.5 8.8 6.4 3.8 3.5 6.9 4.1 5.7 7.4 6.9 6.2 5.3 5.1 4.8 3.6 5.5 3 4 5.4 5.5 3.6 3.5 3.9 4 4 4.9 5.8 5.4 5.5 5.8 5.5 3.7 5.9 3.7 3.5 3.5 4.3 5.8 3.9 3.2 4.6 3.7 5.2 4.3 4.8 4.5 4.3 5.6 4.7 4.5 5.7 6.2 5 4.3 5.7 4.6 2.1 4.6 5.2 4.5 5.2 7 4.2 6.5 3.7 5.9 7 5.8 7.2 5.6 6.5 4.4 1.9 7 3.2 5.5 2.6 6.2 6.2 4.7 6.4 4.2 4.9 2.5 7.4 8.1 5.5 5.1 5.7 4.8 8.2 6.6 5.9 6 7.8 7.3 7.3 7 5 5.7 7.4 4.4 4.2 7.2 7.7 7.2 7.3 6.4 6.3 6.6 8.2 6.7 4.8 3.9 8 5.1 6.6 6.3 6.5 6.3 6.6 8.1 6.6 6.4 6.6 7.7 5.4 7.5 4.4 5.1 2.8 5.2 4.3 4.2 4.6 6.3 5.6 3.5 4 6.1 5.6 4.2 5.9 5.5 5.7 3.5 4 2.4 2.9 4.3 2.1 2.6 4 4.7 3.4 4.2 4.7 4.2 3.4 4.8 2.3 4.1 5.3 4 2.8 4.2 3.8 4.3 4.9 4.1 4.3 4.7 4.1 3.7 4 4.7 4.6 4.8 3 5.3 3.9 4.2 3.5 5.4 4.6 5.2 5.8 4.5 5.6 3.9 5.6 3.4 2.6 4.5 4.5 4.5 4.9 6.7 6.6 4.7 4.9 3.4 5.1 4.3 2.7 5.3 2.9 2.9 5.1 6.2 3.8 4.4 3.8 2.4 5.4 2.6 6.4 4.8 4.4 4.4 3.7 4.3 3.1 4.8 3 4.7 4.2 5.7 4.5 4.4 4.1 3.5 4.5 3.7 4.5 2.3 4.6 4.4 4.9 5.2 6.4 6.9 6.8 5.9 4.5 8.1 5 4 6.1 5.7 4.9 4.8 7.1 4.7 3.4 3.4 7.8 5.1 6.5 4.4 5.8 8.2 6.8 6.1 7.8 7.5 4 5 6.8 7.1
230 227 226 214 225 220 230 222 225 223 227 226 230 223 225 228 215 228 222 206 228 229 222 232 232 236 224 224 229 240 224 222 231 226 221 233 227 232 224 223 223 223 229 223 227 234 227 236 225 232 234 230 235 232 243 237 235 250 236 244 245 234 251 239 241 235 223 248 239 242 244 238 229 249 236 240 237 229 244 254 245 237 238 234 229 235 229 220 226 233 230 234 226 243 234 240 225 233 244 241 238 245 243 248 239 258 252 247 240 249 239 243 244 245 239 241 239 237 232 244 236 241 239 240 233 231 237 229 224 222 237 225 234 228 231 241 246 234 241 230 247 238 231 229 244 243 223 239 245 238 234 249 241 249 240 255 244 244 247 242 245 246 246 248 242 242 245 245 251 240 248 247 247 243 241 247 245 253 251 247 249 243 245 257 245 245 241 244 236 244 244 249 246 250 251 247 240 240 239 251 245 234 238 240 232 246 237 240 233 242 229 236 227 239 236 231 229 235 241 232 235 235 232 231 241 245 219 248 240 234 233 236 252 242 251 224 237 244 250 240 245 243 241 249 226 244 239 233 233 236 234 240 229 235 233 220 234 230 227 226 225 225 224 218 232 229 221 223 235 238 235 235 238 232 231 238 228 237 232 241 232 240 236 240 235 239 229 239 229 248 237 236 243 231 227 231 228 215 227 217 209 212 224 235 224 224 222 231 220 243 234 230 232 225 227 228 233 222 240 223 242 236 236 245 251 244 245 257 245 267 251 257 250 243 258 254 246 248 251 241 253 256 244 245 250
12.9 12.8 12.7 12.6 12.6 12.4 12.3 12.2 12.2 12.1 12.1 12 11.9 11.9 12 12 11.8 11.7 11.6 11.6 11.5 11.4 11.3 11.3 11.2 11.2 11.1 11.1 11.1 11 11 11 11 10.9 10.9 10.8 10.8 10.8 10.8 10.7 10.7 10.6 10.6 10.5 10.6 10.6 10.8 11 11.3 11.4 11.4 11.4 11.4 11.5 11.7 11.9 12 12.1 12.2 12.3 12.2 12.4 12.5 12.4 12.6 12.6 12.6 12.7 12.7 12.7 12.7 12.8 12.8 12.7 12.9 12.8 12.9 12.9 12.8 12.8 12.7 12.7 12.7 12.7 12.6 12.5 12.4 12.3 12.3 12.2 12.1 12.1 12.1 12 11.9 12 12 11.9 11.9 11.8 11.8 11.8 11.8 11.8 11.8 11.8 11.8 11.7 11.7 11.7 11.7 11.6 11.6 11.6 11.5 11.5 11.5 11.6 11.6 11.6 11.7 11.7 11.7 11.7 11.6 11.6 11.7 11.7 11.7 11.6 11.6 11.6 11.6 11.5 11.5 11.6 11.6 11.6 11.6 11.6 11.5 11.5 11.5 11.5 11.6 11.7 11.6 11.5 11.5 11.5 11.6 11.6 11.6 11.6 11.7 11.7 11.9 11.8 11.7 11.7 11.8 11.9 11.9 11.8 11.7 11.8 11.8 11.8 11.8 11.8 11.8 11.8 11.7 11.7 11.6 11.6 11.7 11.6 11.6 11.6 11.7 11.7 11.7 11.6 11.6 11.6 11.5 11.5 11.5 11.5 11.6 11.5 11.5 11.5 11.6 11.6 11.6 11.6 11.6 11.6 11.6 11.6 11.6 11.5 11.5 11.4 11.3 11.2 11.1 11.1 11.2 11.2 11.2 11.1 11.1 11.1 11 11 11 10.8 10.8 10.7 10.7 10.6 10.5 10.5 10.5 10.4 10.4 10.4 10.5 10.6 10.5 10.6 10.6 10.6 10.5 10.5 10.6 10.5 10.5 10.4 10.4 10.4 10.4 10.3 10.3 10.2 10.2 10.2 10.2 10.1 10.1 10.1 10 10.1 10 10.1 10 10 10 10 9.9 10 9.9 9.9 9.9 9.9 9.8 9.8 9.7 9.7 9.8 9.8 9.9 9.8 9.7 9.7 9.7 9.6 9.6 9.6 9.6 9.5 9.4 9.5 9.6 9.7 9.7 9.7 9.7 9.7 9.9 9.8 9.7 9.8 9.7 9.6 9.6 9.7 9.7 9.7 9.7 9.8 9.8 9.8 9.7 9.6 9.6 9.5 9.6 9.6 9.7 9.7 9.8 9.9 10 10 10.1 10.2 10.4 10.5 10.4 10.5 10.6 10.7 10.7 10.8 10.8 10.9 11 11.1 11.2 11.2 11.3 11.3 11.3 11.4 11.5 11.5 11.5 11.5 11.5 11.6 11.6
22.1 21 20.2 19.6 19 18.9 19.8 20.7 21.3 21.6 21.8 22 22.1 22.2 22.3 22.4 22.5 22.7 22.7 22.2 22 22.2 21.5 20.5 19.7 19.1 18.8 19.6 20.6 21.1 21.5 21.7 21.9 22.1 22.2 22.3 22.4 22.5 22.5 22.6 22.4 21.9 22 21.9 20.9 20 19.3 18.8 19.3 20.3 21 21.4 21.7 21.8 22 22.2 22.3 22.3 22.4 22.5 22.6 22.6 22.2 21.9 22.1 21.6 20.5 19.7 19.2 18.9 19.6 20.6 21.2 21.5 21.8 21.9 22.1 22.1 22.3 22.4 22.5 22.5 22.6 22.5 22 22 22.1 21.1 20.2 19.5 19 19 20.1 20.8 21.3 21.6 21.8 21.9 22 22.2 22.3 22.4 22.5 22.6 22.6 22.3 22 22.1 21.8 20.6 19.8 19.3 18.8 19.3 20.3 21 21.4 21.6 21.8 22 22.1 22.2 22.3 22.4 22.5 22.5 22.6 22.3 21.9 22.1 21.7 20.6 19.8 19.2 18.8 19.4 20.3 20.9 21.3 21.6 21.8 21.9 22.1 22.2 22.2 22.3 22.4 22.5 22.6 22.3 21.8 22 21.9 20.8 19.9 19.2 18.8 19.4 20.4 21 21.3 21.6 21.7 21.9 22 22.1 22.2 22.4 22.4 22.5 22.6 22.5 21.9 22 22.1 21 20 19.4 18.9 18.9 19.8 20.7 21.2 21.5 21.7 21.8 21.9 22.1 22.2 22.2 22.3 22.4 22.5 22.6 22.3 21.9 22 21.8 20.7 19.8 19.2 18.7 19.5 20.4 21 21.4 21.6 21.8 21.9 22 22.1 22.2 22.3 22.4 22.4 22.5 22.5 21.9 21.9 22 21.1 20.2 19.3 18.9 18.9 20 20.8 21.2 21.5 21.6 21.8 21.9 22 22.1 22.2 22.3 22.4 22.4 22.5 22.5 21.9 21.9 22 21 20 19.3 18.8 19 20 20.7 21.1 21.4 21.6 21.8 21.9 22 22.1 22.2 22.2 22.3 22.4 22.5 22.5 22 21.8 22 21.3 20.2 19.4 18.9 19 20 20.7 21.2 21.5 21.6 21.8 21.9 22 22.1 22.2 22.3 22.4 22.4 22.4 22.5 22.2 21.8 22 21.5 20.3 19.4 18.9 18.7 19.7 20.5 21 21.3 21.5 21.7 21.8 21.9 22.1 22.2 22.3 22.3 22.3 22.5 22.5 22.4 21.9 21.9 22 21 20.1 19.3 18.7 19.4 20.3 21 21.3 21.6 21.7 21.9 22 22.1 22.2 22.3 22.4 22.5 22.6 22.6 22.4 21.9 22 21.9 20.8 19.9 19.3 18.8 19.4 20.3 20.9 21.3 21.6
54 54 55 54 55 56 56 56 57 57 57 57 58 58 57 58 58 59 59 59 60 60 60 61 61 61 61 61 62 62 62 62 62 62 63 63 63 63 63 64 64 64 64 64 64 64 63 62 61 61 61 61 61 60 59 58 58 58 57 57 57 56 57 56 56 56 55 55 55 55 55 54 55 54 54 54 54 54 54 54 54 54 54 54 55 55 56 56 56 57 57 57 57 58 57 57 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 59 59 59 59 59 59 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 58 59 58 59 58 58 58 58 59 58 58 58 58 58 58 58 58 57 58 58 58 57 57 57 58 58 58 58 57 57 57 57 58 58 58 58 58 58 58 58 58 58 58 58 58 59 59 59 59 59 59 59 59 59 59 59 59 59 59 58 58 58 58 59 59 59 60 60 60 61 60 60 60 60 61 61 61 61 61 61 62 62 62 62 63 63 63 64 64 64 63 63 63 63 63 63 63 63 63 63 63 63 64 64 63 63 64 64 64 64 64 65 65 65 65 65 65 65 65 65 65 65 66 66 65 66 66 66 66 67 67 67 67 67 67 66 67 67 67 67 68 67 67 67 68 68 68 67 67 67 67 67 67 66 66 67 66 66 67 67 67 67 67 67 66 66 66 67 67 67 67 67 67 67 67 66 66 66 65 65 65 64 64 64 64 63 63 63 63 63 62 62 62 61 61 61 61 60 60 60 60 60 61 61 61 60
0.41 0.41 0.41 0.44 0.41 0.35 0.34 0.4 0.44 0.4 0.39 0.39 0.36 0.37 0.39 0.43 0.44 0.41 0.42 0.44 0.43 0.46 0.48 0.47 0.48 0.48 0.46 0.41 0.36 0.41 0.41 0.42 0.49 0.47 0.44 0.46 0.5 0.51 0.48 0.44 0.45 0.46 0.44 0.47 0.47 0.47 0.48 0.45 0.42 0.41 0.48 0.47 0.44 0.45 0.45 0.48 0.52 0.52 0.53 0.53 0.49 0.5 0.52 0.52 0.51 0.53 0.58 0.55 0.51 0.5 0.47 0.48 0.5 0.45 0.45 0.49 0.5 0.47 0.46 0.47 0.48 0.46 0.4 0.4 0.44 0.47 0.5 0.47 0.44 0.46 0.46 0.42 0.4 0.46 0.48 0.48 0.48 0.47 0.49 0.44 0.42 0.46 0.44 0.38 0.39 0.45 0.47 0.45 0.44 0.5 0.52 0.5 0.49 0.5 0.5 0.51 0.5 0.51 0.49 0.5 0.53 0.55 0.51 0.47 0.45 0.44 0.44 0.44 0.45 0.44 0.41 0.42 0.47 0.5 0.49 0.53 0.53 0.48 0.45 0.45 0.48 0.51 0.47 0.48 0.51 0.47 0.44 0.44 0.47 0.49 0.5 0.47 0.48 0.51 0.52 0.48 0.46 0.46 0.47 0.48 0.52 0.55 0.51 0.49 0.49 0.52 0.56 0.52 0.52 0.55 0.55 0.56 0.6 0.62 0.58 0.57 0.6 0.62 0.63 0.6 0.57 0.59 0.64 0.65 0.64 0.61 0.6 0.6 0.65 0.69 0.62 0.59 0.62 0.66 0.68 0.59 0.57 0.63 0.63 0.59 0.59 0.63 0.62 0.61 0.63 0.62 0.63 0.65 0.59 0.59 0.63 0.6 0.61 0.64 0.61 0.61 0.64 0.63 0.58 0.57 0.58 0.59 0.61 0.64 0.62 0.56 0.59 0.62 0.56 0.54 0.54 0.53 0.59 0.62 0.62 0.6 0.56 0.58 0.62 0.6 0.56 0.56 0.6 0.58 0.58 0.59 0.52 0.52 0.6 0.63 0.57 0.51 0.51 0.56 0.57 0.58 0.61 0.62 0.64 0.63 0.6 0.62 0.64 0.63 0.62 0.64 0.64 0.65 0.66 0.63 0.63 0.66 0.66 0.65 0.68 0.69 0.63 0.57 0.56 0.59 0.54 0.5 0.46 0.52 0.58 0.5 0.43 0.45 0.48 0.52 0.51 0.46 0.43 0.43 0.41 0.43 0.48 0.5 0.52 0.47 Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib
0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 1 1 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 1 1 1 1 0 0 0 0 0 0 0 0 0 0 0 0 0 0 1 1 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 1 1 1 1 1 1 1 1 1 1 1 1 1 1 1 1 1 1 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 1 0 0 0 0 0 0 0 0 0 0 1 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 0 1 1 1 1 1 1 1 1 1 1 1 1 1 1 1 1 0 0 0 0 0 0 0 0 0 1 1 Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib
1.96 1.96 1.96 1.97 1.96 1.96 1.96 1.96 1.96 1.97 1.95 1.95 1.96 1.96 1.97 1.96 1.96 1.96 1.96 1.96 1.97 1.96 1.97 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.95 1.95 1.96 1.96 1.96 1.95 1.96 1.95 1.95 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.97 1.96 1.97 1.96 1.96 1.96 1.97 1.97 1.97 1.97 1.97 1.96 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.96 1.97 1.97 1.97 1.97 1.96 1.96 1.96 1.97 1.97 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.97 1.97 1.96 1.96 1.97 1.96 1.96 1.97 1.96 1.96 1.96 1.96 1.96 1.96 1.97 1.96 1.96 1.96 1.96 1.97 1.96 1.97 1.96 1.96 1.96 1.97 1.96 1.97 1.96 1.96 1.96 1.97 1.97 1.96 1.97 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.97 1.97 1.96 1.96 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.98 1.97 1.98 1.97 1.97 1.97 1.97 1.98 1.98 1.97 1.97 1.98 1.98 1.98 1.98 1.97 1.97 1.97 1.97 1.97 1.97 1.98 1.97 1.97 1.97 1.98 1.98 1.98 1.97 1.98 1.98 1.97 1.97 1.97 1.98 1.97 1.97 1.98 1.98 1.98 1.98 1.97 1.97 1.97 1.98 1.98 1.97 1.97 1.97 1.98 1.98 1.97 1.97 1.97 1.98 1.97 1.97 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.99 1.98 1.98 1.98 1.96 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.99 1.98 1.98 1.98 1.99 1.99 1.98 1.98 1.98 1.98 1.98 1.99 1.97 1.98 1.98 1.98 1.97 1.98 1.97 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.99 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.98 1.97 1.97 1.97 1.97 1.97 1.97 1.98 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib
1.96 1.95 1.95 1.96 1.96 1.95 1.95 1.95 1.96 1.96 1.95 1.94 1.95 1.96 1.96 1.96 1.95 1.96 1.95 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.95 1.96 1.96 1.96 1.95 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.97 1.96 1.96 1.96 1.96 1.96 1.97 1.96 1.96 1.96 1.96 1.96 1.96 1.95 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.95 1.95 1.95 1.95 1.96 1.96 1.96 1.96 1.96 1.95 1.96 1.96 1.95 1.95 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.95 1.95 1.94 1.96 1.95 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.95 1.95 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.96 1.97 1.97 1.97 1.97 1.97 1.97 1.96 1.96 1.97 1.96 1.96 1.96 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.96 1.96 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.96 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.96 1.96 1.97 1.97 1.97 1.97 1.96 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.98 1.97 1.97 1.97 1.95 1.97 1.97 1.98 1.98 1.98 1.97 1.97 1.98 1.97 1.98 1.98 1.98 1.98 1.97 1.97 1.97 1.98 1.98 1.97 1.97 1.97 1.97 1.97 1.98 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.98 1.97 1.97 1.98 1.98 1.97 1.97 1.97 1.98 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.97 1.96 1.96 1.96 1.97 1.96 1.97 1.96 1.96 1.96 1.96 1.97 1.96 1.96 1.96 Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib Calib
Chart: THC

## Slide 9
 | Reviews
WS/WD | If hourly data do not change a lot or if WS data is recorded <1.0kph for more than 5 hours => sensor failure
RH/BP/TPX (TPX: Temperature) | Hourly values are consistent for more than 5 hours => sensor failure Rule of Thumb: TPX and RH are generally inversely related.
STNTPX | STNTPX ideally should be kept in a range of 20 to 25 degC Temperature changes should not be more than 5 degC in a day.

## Slide 10: Wind System Failure Low Shelter TPX
Notes: These are examples of the wind system failure and low shelter temperature.

## Slide 11
Diff (Column K): the difference between the Ref ( column H) and the Meas (measured) (column I)
e.g. Zero drift= (-0.04)-0=-0.04
Span drift=(44.81-44.65)/44.65*100
= 0.36%
Diff Last Cal (Column L): the ‘’Diff’’ value from the previous day
Diff From Last (Column M): the difference between the values in Columns K and L

## Slide 12: TRS failed span check
NMHC span check failed