## Slide 1
‘Daily’ Zero / Span Checks
What?
  A check of the analyzer’s accuracy
  Uses an ‘internal’ source of zero air
  Uses an ‘internal’ source of test gas
  Qualitative not quantitative!

## Slide 2
‘Daily’ Zero / Span Checks
Why?
Quality Control
Demonstrates that analyzer continues to work as intended
Identifies potential equipment failure
Can be used to determine ‘point of failure’

## Slide 3
‘Daily’ Zero / Span Checks
How?
Analyzer
Detector
Exhaust
Sample
Zero Air source
Valve
Test gas source

## Slide 4
‘Daily’ Zero / Span Checks
Zero Air Source:
Compressed gas (cylinder)
Scrubber
Zero Air Generator

## Slide 5
‘Daily’ Zero / Span Checks
Test Gas Source:
Compressed gas (cylinder) Permeation (Perm) Tube and Oven

## Slide 6
‘Daily’ Zero / Span Checks
When?
  After multi-point calibration
  After span gas exchange
    Sets Expected (reference) value
  Daily
    Zero ± 3% of full scale; Span ± 10% of reference (AMD)
  Unscheduled (repeat) if suspected problem
  To bracket ‘events’ (e.g. power)
    Demonstrates good recovery